Protein Assay – Bio-Rad

Jennie Baier with alterations from Beth Furnish 8/99 Protocol and Bio-Rad Product Information Sheets

UT, 6/03

### Preparation

Bio-Rad Assay Kit with BSA (Bovine Serum Albumin) Standards (cat #500-0002)

Read MSDS--contains methanol and butyl acetate and phosphoric acid!

Plate reader with 595 nm filter

96 well plate, multi-channel or single channel pipetmen and tips

Dye Reagent:

4 ml	Dye Reagent	(1 part)

16 ml	ddI		(4 parts)

20 ml	volume necessary for one 96 well plate assay

The solution is stable for about 2 weeks at room temperature. Filter through Whatman #1 filter if necessary.

## Standard Procedure: 50-500 μg/ml detection range

- Prepare dilutions of standard protein. BIORAD supplies BSA standard at 1.4 mg/ml in ddI. Standard solutions are stable at 4(C for 2 months. The following chart prepares 0.5 ml of each standard (enough for about 15 assays).

Standard Number	Standard Conc. (mg/ml)	BSA Solution (μl)	ddI (μl)

S1	0.05 	17.9	482.1

S2	0.15	53.6	446.4

S3	0.25	89.3	410.7

S4	0.35	125.0	375.0

#### S5	0.50	178.6	321.4

- Pipet 10 μl of each standard and sample into the wells. Use ddI or buffer for the blanks.

- Pipet 200 μl of diluted dye reagent into each well and mix on an orbital or platform shaker.

- Incubate 5 minutes to 45 minutes (no more than 1 hour) and measure absorbance at 595 nm.

## Microassay Procedure: 8-80 μg/ml detection range

- Prepare dilutions of standard protein. BIORAD supplies BSA standard at 1.37 mg/ml in ddI. Standard solutions are stable at 4(C for 2 months. The following chart prepares 1.0 ml of each standard.

Standard Number	Standard Conc. (μg/ml)	BSA Solution (μl)	ddI (μl)

M1	1.0	0.73	999.3

M2	5.0	3.65	996.4

M3	8.0	5.84	994.2

M4	20.0	14.6	985.4

M5	40.0	29.2	970.8

M6	60.0	43.8	956.2

#### M7	80.0	58.39	941.6

- Pipet 160 μl of each standard and sample into the wells. Use ddI or buffer for the blanks.

- Pipet 40 μl of the dye reagent concentrate into each well and mix on an orbital or platform shaker.

- Incubate 5 minutes to 45 minutes (no more than 1 hour) and measure absorbance at 595 nm.